## Slide 1: Truventic QCL SOP
24 Sep 2020

## Slide 2
Truventic QCL Emission Spectrum
4.45 mm

## Slide 3
Safe operating point
Don’t go here!
Threshold = 1.1 Amp

## Slide 4
Voltage compliance set here at danger zone
Usual operating point

## Slide 5
Power supply for Temperature contoller.
AC connections.
Hot (usually black, here brown) measures 120 VAC to ground.
Return (usually white, here blue) measures zero volts AC to ground (green).
To Temperature Contoller
This PS is kaput. Needs to be replaced.

## Slide 6
Temperature controller. Operating Temperature = 20 oC.
Thermister

## Slide 7
T-control
Short for when laser not in use.
Plug in when laser is to be used.
To light wave laser driver
This thermistor is used with T-controller

## Slide 8
Lightwave laser driver
Make adjustments with output off. Choose 2000 mA range. Set knob to zero. Operate in current mode with limit 2.02 Amp. Set compliance voltage to 11.5 V.
Turn output on
Slowly increase current knob to 2 A